# QDM-ELN-Import — how to use it

Records your DynaCool measurements in the electronic lab notebook, so you do not have to write them up afterwards.

You add two steps to your sequence. At the start, a window appears asking who you are and which sample this is. When the measurement finishes, everything the instrument produced is added to the same entry by itself.

## 1. What you have to do, once per sequence

Add a Call Script step at the beginning and at the end of your sequence. Both point at the same file:

New Datafile "C:\QdDynacool\Data\MySample.dat"

Call Script C:\QdDynacool\QDM-ELN-Import\QDM-ELN-Import.bas      ← near the start

...  your measurement steps  ...

Call Script C:\QdDynacool\QDM-ELN-Import\QDM-ELN-Import.bas      ← last step

When you add the step, change the file type in the open dialog. It starts on *Sequence Files (\*.seq)* and the script simply will not be listed. Switch to SAX Basic Script Files (\*.bas), then pick QDM-ELN-Import.bas.  This catches everybody at least once — it looks as though the file is missing.

Both steps must be the same file. MultiVu cannot run two different scripts in one sequence. The program works out by itself which call is the first and which is the last.

### Where to put the first call

Put it just after the first `New Datafile` step if you can. Then the data file already exists, its header can be read, and your sample is usually recognised and selected before the window even opens.

Earlier is not wrong — nothing is lost, you just search for the sample yourself.

### The last call is optional, but do add it

Without it the entry is still created; it just will not contain the data-file details, and it keeps a note saying the measurement did not finish.

## 2. What happens at the start

A window opens. The sequence waits for it, so fill it in and press *Send to Wiki*.

If you walk away, it gives up by itself after ten minutes and the measurement continues without an entry — you can add it by hand later.

Fields marked with a red star must be filled in.

### Context — who and what

Operator. Start typing your name; matches from the staff directory appear below, and the resolved name is shown in bold. Your last entry is remembered, so usually it is already correct.

Or scan your HZDR ID card. Its QR code sets the operator to you, and it wins over everything else — in particular over the person who *prepared* the sample, who is often not the person running the measurement.  The order does not matter. Card first then sample, or sample first then card: the result is the same. A sample label always contains a UUID and an ID card does not, so the program can always tell which you scanned.  Once your card is scanned the operator field stops asking you to confirm, because the question is answered.

Project and Sample series are chosen from lists read from the wiki. If something is missing, the button Add or edit Sample, Project or Sample Series to Wiki opens the wiki so you can create it, then reopen the sequence.

If the operator or sample name blinks, look at it. It means either a new sample was loaded, or the remembered values are more than three hours old and probably belong to somebody else.

### Sample — which sample was measured

Three ways, easiest first:

1. Scan the QR code on the sample label with the barcode scanner. The sample is looked up and

selected immediately. Just scan — no field needs to be clicked first.

2. The instrument already knows. If you typed a sample name into the VSM or ETO dialog, it is

searched for automatically. One match is taken; several make the match list blink; none makes the sample name blink.

3. Type part of the name and press *Search*. Narrow the list with *only my samples*,

*this project* or *this sample series*.

The selected sample is shown in bold. A measurement cannot be stored without one — that is what the whole record hangs on.

If a scan is refused, the message says why: the label belongs to the old Athene wiki, the sample does not exist in this wiki, or the code is not a sample address at all. The offending address is always shown so you can see what was scanned.

### Measurement, Conditions, Data files

Read from your sequence and normally correct already:

• Option — VSM or ETO.

• Type — M(H), M(T), R(T), IV(H) and so on. If your sequence does something the

program cannot classify it says custom, and then a comment is required.

• Field / temperature ranges and step sizes, and the data files the sequence writes.

• Initial (virgin) curve measured — ticked when the measurement started at zero field.

Correct anything that is wrong. What you type always wins over what was detected.

### Comments

Free text.

Context written at the instrument is picked up automatically. Put either or both of these in the sample comment (VSM or ETO sample properties), or as a REM line in the sequence:

Project = [ThickFilmSputterDeposition]

SampleSeries = [SebastiansTestSeries]

Both on one line works too: Project = [ThickFilmSputterDeposition] / SampleSeries = [SebastiansTestSeries]

They preselect the Project and Sample series boxes — useful when you measure for an external group, so the context is typed once rather than every run. Scanning a sample still overrules them, and so does anything you choose yourself.

Put the value in square brackets. They mark where the name ends, so you can write anything else on the same line:  | written | taken as | |---|---| | Project = [CoFe] remeasured after annealing | CoFe | | Project = CoFe remeasured after annealing | CoFe remeasured after annealing ✗ |  Without brackets the name has to run to the end of the line, because nothing says where it stops. It still works — Project = CoFe on its own is fine — but the brackets let you keep your notes and the context on the same line.

Either the plain name or the short identifier works. Project = [Thick film sputter deposition] and Project = [ThickFilmSputterDeposition] both find the same project. The identifier is the more durable of the two, since a description can be reworded later and the name then stops matching — but use whichever you find easier to remember. Both are shown in the dialog's Project and Sample series boxes.

It must be written before the run starts, because it is read when the sequence begins. Typing it into this Comments box has no effect on those two boxes — by then they are already filled in.

## 3. What happens at the end

Nothing you need to do. A small window shows what was added and closes itself after a few seconds.

Added automatically: the data files the run produced, the field and temperature actually reached, the chamber pressure, averaging time, oscillation amplitude, excitation current and frequency, and the instrument serial numbers.

Nothing you typed is overwritten — only fields that were still empty get filled.

If you ticked Email me when the measurement is finished, you now get a mail with the sequence name and a link to the entry. It is sent from your own address, so it appears in your mailbox as if you had written the reminder yourself.

## 4. Reading the entry later

The entry is a page in the wiki, linked from the sample. It carries everything above, plus the complete sequence as MultiVu rendered it — so it can be reproduced.

An entry that still shows `QDM-ELN-Import: Starting…` in its comments belongs to a run that never reached its final step: aborted, or the last Call Script was missing. The measurement may be perfectly good; the entry is simply incomplete.

## 5. If something does not work

| What you see | What it means |
| --- | --- |
| The script is not in the file list | The file type is still *Sequence Files (\*.seq)* — switch to *SAX Basic Script Files (\*.bas)* (§1) |
| No window appears | The Call Script step points somewhere else, or the program is not installed on this machine |
| "COM server not available" | The program is not registered. Ask whoever installed it |
| The scanner beeps but nothing happens | Wrong USB socket or wrong driver — a known problem, ask for help |
| "Sample not existing in this wiki" | The label is genuine but the sample was never created. Use the wiki button and add it |
| "Invalid URL" | What was scanned is not a sample address at all |
| No entry appeared | Look in C:\QdDynacool\QDM-ELN-Import\logs\last-problem.txt — it says what went wrong in plain words |

Your measurement is never at risk. The program cannot start, change or stop a measurement; it only reads and writes records. If it fails, the run continues and the data file is written exactly as always. The data file is the authoritative record — the wiki entry is a convenience.

## 6. Good habits

• Give sequences unique names. If two folders hold a sequence with the same name, you will be

asked which one you used. Adding your initials avoids the question for good.

• Set the data file in the sequence, not only in the VSM dialog. A New Datafile step

overwrites what you selected there — and the change stays after the run.

• Keep sequences on a shared drive, so colleagues can open what the entry points at.

• Scan the label rather than typing a name. It cannot select the wrong sample.